## MPPC ETD
MOFE on Board TAD, PT 1000 connections and variable bias functionality
TAD Bias functionality flow:
|  | 1) | Read MPPC Vop: MPPC_ETD_0_VopCal
In the case of failure EEPROM reading generate message: Detector EEPROM cannot be read. Verify detector connection.
During the Bias test, set bias voltage cannot exceed value of: MPPC_ETD_0_VopCal.
|  | 2) | Using the MOFE calibration data bias voltage is set in N steps to VopCal and is alse read using OD TAD.
Reading using OD TAD is done as:
Set:
MOFE_0_TADSwitch = 1 (On)
Get:
MOFE_0_CtrlCheck
How the bias is set is illustrated in the attached file. Note that I use just default MOFE calibration data, in reality real calibration data is read from OD.
The result is plot this graph:
[5 empty] | x-axis |  | MOFE_0_Bias
[5 empty] | y-axis |  | MOFE_CtrlCheck
To pass this TAD, for each point the allowed difference is 10%
For more details illustration script is attached:
# -*- coding: utf-8 -*-
" Created on Wed Jun 5 14:42:39 2019 @author: petr.kucera "
import math
import numpy as np
from scipy import misc
from scipy import signal
import matplotlib
import matplotlib.pyplot as plt
def RangeConversion(Values_in,min_in,max_in,min_out,max_out):
# linear function
k = (max_out-min_out)/(max_in-min_in)
c = min_out - k*min_in
Values_out = k*Values_in+c
return Values_out
# MOFE calibration values read from MOFE
MOFE_0_Bias0Cal = 0
MOFE_0_BiasRangeThrLoCal = 25
MOFE_0_BiasRangeThrHiCal = 50
MOFE_0_MaxBiasOffsetCal = 30
# Reading TAD object
# MOFE_0_CtrlCheck
# MOFE Set Object
# MOFE_0_Bias
# Max_MOFE_0_Bias (form OD definition)
Max_MOFE_0_Bias = 2.045
# Read from MPPC ETD EEPROM
MPPC_ETD_0_VopCal = 55
# only for plotting purpose
Bias_out_array = np.array([])
Range_array = np.array([])
Bias_array = np.linspace(MOFE_0_Bias0Cal,MPPC_ETD_0_VopCal,10)
for Bias in Bias_array:
if (Bias <= MOFE_0_BiasRangeThrLoCal):
# Set bias range to HW
MOFE_0_BiasRange = 0
# Set bias control DAC to HW
MOFE_0_Bias = RangeConversion(Bias-MOFE_0_Bias0Cal,0,MOFE_0_MaxBiasOffsetCal,0,Max_MOFE_0_Bias)
elif (Bias <= MOFE_0_BiasRangeThrHiCal):
MOFE_0_BiasRange = 1
MOFE_0_Bias = RangeConversion(Bias-MOFE_0_BiasRangeThrLoCal,0,MOFE_0_MaxBiasOffsetCal,0,Max_MOFE_0_Bias)
else:
MOFE_0_BiasRange = 2
MOFE_0_Bias = RangeConversion(Bias-MOFE_0_BiasRangeThrHiCal,0,MOFE_0_MaxBiasOffsetCal,0,Max_MOFE_0_Bias)
# only for plotting purpose
Bias_out_array = np.append(Bias_out_array,MOFE_0_Bias)
Range_array = np.append(Range_array,MOFE_0_BiasRange)
# Check based on TAD reading of:
# MOFE_0_CtrlCheck
# if (abs(MOFE_0_CtrlCheck-MOFE_0_Bias)/MOFE_0_Bias) < 0.1:
# print('Bias check passed')
# else:
# print('Bias check failed')
data1 = Bias_out_array
data2 = Range_array
t = Bias_array
fig, ax1 = plt.subplots()
color = 'red'
ax1.set_xlabel("Bias [V]")
ax1.set_ylabel('MOFE_0_Bias [V]', color=color)
ax1.plot(t, data1,'o-',color=color)
ax1.tick_params(axis='y', labelcolor=color)
ax2 = ax1.twinx() # instantiate a second axes that shares the same x-axis
color = 'blue'
ax2.set_ylabel('MOFE_0_BiasRange', color=color) # we already handled the x-label with ax1
ax2.plot(t, data2,'o-',color=color)
ax2.tick_params(axis='y', labelcolor=color)
fig.tight_layout() # otherwise the right y-label is slightly clipped
plt.grid()
plt.show()
PT 1000 connection
The purpose of this TAD is to check if the thermistor PT1000 is properly connected
For instance (again default MOFE calibration values assumed, but real has to be read):
|  | MOFE_0_RefVTempCal |  |  |  | 2.6 | V
|  | MOFE_0_MaxTempSensCal |  |  |  | 110 | mV/Deg
Set:
|  |  | MOFE_0_Temp |  |  | 1.0225 |  | Set to half its OD range
MOFE_0_TADSwitch = 0 (Off)
Get:
|  |  | MOFE_0_TMeasOut |  |  | 1 |  | Example of read value
Calculate temperature as:
Calculate temperature as:
|  |  | T[Deg] = 1000*(MOFE_0_TMeasOut - MOFE_0_RefVTempCal/2)/(MOFE_0_MaxTempSensCal/2) + 25Deg | [7 empty] | 19.5455
|  |  | T[Deg] | 19.5455
Report results
o T < = 5 Deg:
Measured temperature is xxx. MPPC temperature is too low. Temperature resistor seems to be shorted. Replace MPPC ETD.
o T >= 45 Deg
Measured temperature is xxx. MPPC temperature is too high. Temperature resistor seems to be disconnected. Replace MPPC ETD.
o 5 Deg< T<45 Deg;
Measured temperature is xxx.
MPPC ETD System TAD test
Purpose of this part is troubleshoot overall MPPC ETD signal path from different point of view then boards TAD
ETD - Bias Calibration
Reasons of fails:
Bad MOFE calibration data (real Bias is not a monotonic function)
Detector issue (low gain, collection efficiency…)
HW issuse (dmaged MOFE or other board)
Buggs, SW issue (NOT assumed)
Optics issue (badly adjusted spots…), NOT aasumed her
ETD - MOFE offset calibration; modification for TAD
Results of the calibration is function:
MOFE_Offset = f(Bias) to get 0V output
Basically parasitic MOFE offset and MPPC dark current is measured
|  | Steps: |  | (but these valuese can be deduced from offsets obtained in the calibration above)
|  |  | 1) | Determine required MOFE_Offset for 0V bias set
|  |  |  | passed: |  | abs(offset) <= 0.1
|  |  |  | failed: |  | abs(offset) > 0.1 |  |  | MOFE damaged
|  |  | 2) | Determine dark current mean value for bias = Vop |  |  |  |  | (DABE setting as in ETD - MOFE offset calibration)
|  |  |  | passed: |  | within the range 40-80DN
|  |  |  | failed: |  | <40DN |  | Wrong Vop, PT1000 issue, MOFE issue, wrong MOFE calibration data, Detector issue
[6 empty] | >80DN |  | Wrong Vop, PT1000 issue, MOFE issue, wrong MOFE calibration data, Detector issue
was measured for 20dB on DABE and PreOn
MOFE calibration uses 10dB
Vop was about 54.8V
Do 60DN id a typical value for DABE gain 10dB and PreAmp ON
MOFE calibration data check; range consistency
With beam and signal acquistion
Vop cannot be exceeded
At points where there are changes in the discrete MOFE bias ranges
That bias, where the range changes will be set from left and from right
Same signal has to be measured
MOFE HW or MPPC HW issue
Signal acquisition without beam
Set DABE gain 10dB and PreAmp On
Set Bias 0V
|  |  | White signal or black signal: |  |  | MOFE, DABE, PIA Issue
If ok continue
Set Bias Vop - 15V
Calculate signal difference for setting Bias = (Vop -15V) and Signal 0V
|  |  | if > 50 DN |  | MPPC ETD issue

## TestOnA7
Bias functionality flow
| Read | GET
0x2241 MPPC_ETD_0_VopCal
0x2321 MOFE_0_Bias0Cal
0x2322 MOFE_0_BiasRangeThrLoCal
0x2323 MOFE_0_BiasRangeThrHiCal
0x2324 MOFE_0_MaxBiasOffsetCal
| Tests | SET | GET
| 0x8052 MOFE_0_TADSwitch | 1
| 0x230E MOFE_0_BiasRange | 0
| 0x230F MOFE_0_Bias | 0
0x8054 MOFE_0_CtrlCheck
Test formula:
| 0x230E MOFE_0_BiasRange | 0
| 0x230F MOFE_0_Bias | 0.37492
0x8054 MOFE_0_CtrlCheck
| Test formula: | 1
| 0x230E MOFE_0_BiasRange | 1
| 0x230F MOFE_0_Bias | 0.92026
0x8054 MOFE_0_CtrlCheck
| Test formula: | 1
| 0x230E MOFE_0_BiasRange | 1
| 0x230F MOFE_0_Bias | 1.6701
0x8054 MOFE_0_CtrlCheck
| Test formula: | 1
PT 1000 connection |  | GET | SET
0x2326 MOFE_0_RefVTempCal
0x2325 MOFE_0_MaxTempSensCal
| 0x2310 MOFE_0_Temp |  | 1.0225
| 0x8052 MOFE_0_TADSwitch |  | 0
0x8053 MOFE_0_TMeasOut

## MOFE Cal Params Expl
MPPC parameters :
Temperature coefficient (TemSens):
54 mV/Deg; saved at the detector EEPROM
TempSens1 = 54
Gain at 25Deg: 1.7 E6; (saved at the detector EEPROM but not used)
Vop: saved at the detector EEPROM, unique for each MPPC
When Vop is set as the bias voltage, and temperature of the diode is 25 Deg, then the gain is 1.7E6.
If the temperature is changed to, say, 26Deg then you have to set Vop + 54mV to keep the gain 1.7E6.
MPPC bias range:
We also use bias < Vop to reduce MPPC gain, when the beam current is large.
Let us say Vop = 55V and gain 1.7E6.
We set 54 V and gain is 0.7E6 (we have calibration for this, which measures function: Gain = f(Bias)).
Also for this bias we assume temperature coefficient 54mV/Deg. (But strictly speaking, this temperature coefficient is defined for Vop)
MOFE parameters
Calibration Ods names: | Default value | Units | MOFE#2 | Meanings
MOFE_0_Bias0Cal | 0 | V | -1.2679 | Minimum voltage at range 0 (MOFE_0_BiasRange = 0 and MOFE_0_Bias = 0)
MOFE_0_BiasRangeThrLoCal | 25 | V | 24.018 | Minimum voltage at range 1 (MOFE_0_BiasRange = 1 and MOFE_0_Bias = 0)
MOFE_0_BiasRangeThrHiCal | 50 | V | 48.768 | Minimum voltage at range 2 (MOFE_0_BiasRange = 2 and MOFE_0_Bias = 0)
MOFE_0_MaxBiasOffsetCal | 30 | V | 30.702 | HV Bias span (measured voltage difference for: (MOFE_0_Bias = max) - (MOFE_0_Bias = min))
MOFE_0_MaxTempSensCal | 110 | mV/Deg | 1000*(B27-B17)/10 | Set: MOFE_0_Bias = 0V, MOFE_0_BiasRange = 3, MOFE_0_Temp = max, Output bias is measured for two set temperatures (35 and 25 Deg). Results is: (V35-V25)/(35-25)
MOFE_0_RefVTempCal | 2.6 | V | 0 | Termistor at 25Deg, voltage difference for: (MOFE_0_Temp = max) - (MOFE_0_Temp = min)
MOFE_0_OffsetCal | 0 | V | -20.7mV | Parasitive offset value
MOFE_0_MaxTempSensCal: Measured as:
Temperature control set to Max, i.e. OD MOFE_0_Temp = max (2.045 in FW)
Temperature was changed by 10 Degrees and output bias measured:
1000*(Bias_35Deg-Bias_25Deg)/10Deg
1000 is because it is scaled to mV
MOFE_0_RefVTempCal
Termistor at 25Deg, output voltage difference for: (MOFE_0_Temp = max) - (MOFE_0_Temp = min)
Say MOFE_0_RefVTempCal = 2.6
SW control
TempSens1 = 54 (from above)
// pavel.stejskal@fei.com
// modifications by Petr Kucera
OffsetCmp1 = 0;
if (BiasEnable)
{
double biasOut = 0;
// BiasV1 is just demanded bias value to have particular gain (this used function Gain = f(Bias) got by a calibration, but not important)
// Just say you want to set: BiasV1 = 54V
BiasV1 = FitCoefB1 * pow(10,FitCoefK1 * (GainDb - MaxGainDb)/20) - FitCoefA1 * pow(10,-FitCoefKK1*(GainDb - MaxGainDb)/20) + Vop1 - FitCoefB1 + FitCoefA1;
// Vime, ze kdyz ma dioda teplotu 25Deg, tak teplotni kompenzace pri jejim maximalnim nastaveni pricte: RefVtemp1
// Kdyz je nastaveni teplotni kompenzace TempSens1, tak pro teplotu 25Deg je pricteno: RefVtemp1*TempSens1/MaxTempSens1
// rekneme: 2.6*54/110 = cca 1.3V. Tedy kdyz ma diode teplotu 25Deg a teplotni kompenzace je nastavena na 54mV/Deg tak HW pricte 1.3 V
double corrBias = BiasV1-RefVtemp1*TempSens1/MaxTempSens1;
// corrBias je tedy: pozadovane napeti – napeti_ktere_HW_pricte_kdyz_ma_dioda_25Deg_a_teplotni_kompenzace_je_nastavena_na_54mV
// pokud bude tedy teplota MPPC 25Deg, tak predeslou hodnotu HW zase pricte
// pokud se bude teplota menit, co je rozdilna od 25Deg, tak to HW upravi
// napr. pozadujeme 54V; Tak posleme na HW: 54 – 1.3 = 52.7
// kdyz je teplot 25Deg, tak HW nastavi: 52.7 + 1.3 = 54
// kdyz je teplota 27, tak HW nastavi: 52.7 + 1.3 + 2*54m = 54.108V
// dalsi cast je jen prelozeni 52.7V na FW objekty s pouzitim dalsich kalibrovanych MOFE parametru
if (corrBias <= Bias0Cal1)
{
biasOut = 0;
}
else if (corrBias <= BiasRangeThrLo1)
{
BiasRange1 = 0;
biasOut = RangeConversion(corrBias-Bias0Cal1,0,MaxBiasOffset1,0,2.045);
}
else if (corrBias <= BiasRangeThrHi1)
{
BiasRange1 = 1;
biasOut = RangeConversion(corrBias-BiasRangeThrLo1,0,MaxBiasOffset1,0,2.045);
}
else
{
BiasRange1 = 2;
biasOut = RangeConversion(corrBias-BiasRangeThrHi1,0,MaxBiasOffset1,0,2.045);
}
BiasRaw1 = clamp(biasOut,0,2.045);
// Nastaveni teplotni kompenzace na TempSens = 54mV/Deg
// Kdyz se nastavi FW max hodnota 2.045, tak kompenzace je MaxTempSensCal (cca 110mV)
// pomerove (linearne) se tedy FW hodnota upravi
TempCoefRaw1 = clamp(2.045*TempSens1/MaxTempSens1,0,2.045);
OffsetCmp1 = Offset1;
}
else
{
BiasRange1 = 0;
BiasRaw1 = 0.0;
OffsetCmp1 = 0.0;
TempCoefRaw1 = 0.0;
}
//
if (ChannelNum == 0) Channel0On = BiasEnable;